## 1993-09-22
RECALL_CLASSIFICATION_DATE | PRODUCT_TYPE | CLASSIFICATION | RECALL_NUMBER | PRODUCT | RECALLING_FIRM | MANUFACTURER | RECALL_INITIALIZATION_DATE | REASON | VOLUME | DISTRIBUTION
1993-09-22 | Drugs | III | d-338-3 | Smx-tmp Ss (sulfamethoxazole 400 Mg/trimethoprim 80 Mg Single Strength Tablets) Rx Antibacterial With Indications And Usage For Treatment Of Urinary Tract Infections Acute Otitis Media Chronic Bronchitis And Others. Recall #d-338-3. | Manufacturer | Eon Labs Manufacturing Inc. Laurelton New York. | September 10, 1993 | Overage of the trimethoprim ingredient | Approximately 38 682 Bottles Were Distributed. | Nationwide.
1993-09-22 | Biologics | II | b-392-3 | Recovered Plasma. Recall #b-392-3. | Manufacturer | Lifesource Glenview Illinois. | April 22, 1993 | Blood products collected from a donor with a history of cancer were distributed | 1 Unit. | New York.
1993-09-22 | Biologics | II | b-394-3 | Platelets. Recall #b-394-3. | Manufacturer | American Red Cross Blood Services Charlotte North Carolina. | January 11, 1993 | Blood product, containing anti-c, anti-e, and anti-g antibodies, but labeled negative for unexpected antibodies, was distributed | 1 Unit. | North Carolina.
1993-09-22 | Biologics | II | b-426/428-3 | (a) Red Blood Cells; (b) Platelets; (c) Recovered Plasma. Recall #b-426/428-3. | Manufacturer | American Red Cross Atlanta Georgia. | November 20, 1992 | Blood products which tested non-reactive for the antibody to the human immunodeficiency virus type 1 (anti-hiv-1), but were collected from a donor who previously tested initially reactive for anti-hiv-1, were distributed | 1 Unit Of Each Component. | Georgia California.
1993-09-22 | Biologics | II | b-429-3 | Recovered Plasma. Recall #b-429-3. | Manufacturer | Community Blood Centers Of South Florida Inc. Lauderhill Florida. | March 16, 1993 | Blood product, which tested repeatedly reactive for the antibody to the human immunodeficiency virus type 2 (anti- hiv-2), was distributed | 1 Unit. | Tennessee.
1993-09-22 | Biologics | III | b-391-3 | Red Blood Cells Expired. Recall #b-391-3. | Manufacturer | Lifesource Glenview Illinois. | April 22, 1993 | Blood products collected from a donor with a history of cancer were distributed | 1 Unit. | Tennessee.
1993-09-22 | Biologics | III | b-424-3 | Sebra Tubing Sealers. Recall #b-424-3. | Manufacturer | Engineering And Research Associates Inc. Tucson Arizona. | March 19, 1993 | Blood collection tubing sealers, producing irregular or cut seals, were distributed | 63 Units. | Nationwide And International.
1993-09-22 | Devices | I | z-689-3 | Matrx Lifedefense Plus (ldp) Portable Defibrillator/monitor/pacer Used To Provide Emergency -4- Electric (dc) Defibrillation Ecg Waveform Monitoring And Both Demand And Fixed Transcutaneous Pacing Functions. Recall #z-689-3. | Matrx Medical Inc. Orchard Park New York | Nec Tochigi Ltd. Tochigi Prefecture Japan. | April 21, 1993 | The device will fail to operate in the defibrillation mode if a certain component (the d8 diode located on the high voltage pc board) fails | Approximately 1 385 Units Were Distributed. | Nationwide And International.
1993-09-22 | Devices | I | z-721-3 | Matrx Lifedefense Plus (lpd) Portable Defibrillator/monitor/pacer Used To Provide Emergency Electric (dc) Defibrillation Ecg Waveform Monitoring And Both Demand And Fixed Transcutaneous Pacing Functions. Recall #z-721-3. | Matrx Medical Inc. Orchard Park New York | Nec Tochigi Ltd Tochigi Prefecture Japan. | July 23, 1993 | The relay switch (k1) which turns the unit on could malfunction resulting in the unit failing to turn on, or requiring multiple depression of the on/off button before the unit turns on | 194 Units Were Distributed. | Nationwide And International.
1993-09-22 | Devices | II | z-460/463-3 | (a) Exam-tex Brand Medium/large Gloves Product #7208 (b) Exam-tex Brand Small/medium Gloves Product #7209 (c) Vha+plus Brand Small/medium Gloves Product #v7266 (d) Vha+plus Brand Medium/large Gloves Product #v7267. Recall #z-460/463-3. | Ansell Inc. Dothan Alabama | Ansell Lanka (pvt) Ltd. Byagama Sri Lanka. | March 2, 1993 | Product failed fda leak test | Approximately 100 Million Gloves Were Distributed; Firm Estimated That 10 Million Gloves Remained On The Market At The Time Of Recall. -5- | Nationwide.
1993-09-22 | Devices | II | z-646-3 | Mevatron Linear Accelerator Digital Controlled X-ray And Electron Beam Therapy. Recall #z-646-3. | Manufacturer | Siemens Medical Laboratories Inc. Concord California. -6- | NOT-FOUND | A defective connection link between the target and the target drive assembly | 331 Units Were Distributed. | Nationwide And International.
1993-09-22 | Devices | II | z-654/655-3 | Respiration/heart Rate Monitor Used In The Care Of Patients At Risk From Central Apnea Fast Heart Rate Or Slow Heart Rate And Are Intended To Warn The Care Giver When Breathing Effort Heart Rate Or Both Are Outside Limits Specified By The Physician: (a) Model 9500; (b) Model 9550. Recall #z-654/655-3. | Aequitron Medical Inc. Plymouth Minnesota | Aequitron Medical Inc. Plymouth Minnesota. | June 18, 1993 | Vibration could cause the alarm to malfunction | 178 Units Were Distributed. | Nationwide Germany.
1993-09-22 | Devices | II | z-696\700-3. | Auto Suture Endo Gage 30 Disposable Instrument Reorder #173025 A Sterile Device Used During Laparotomy Procedures Sold Separately Or In Kits: (a) Auto Suture Endo-gauge 30 Disposable Instrument Reorder #173025; (b) Endoscopic Cholecystectomy; (c) Endoscopic Bowel; (d) Endoscopic Hernia; (e) Lavh (laparoscopically Assisted Vaginal Hysterectomy). Recall #z-696\700-3. | U.s. Surgical Corporation North Haven Connecticut | United States Surgical Corporation Norwalk Connecticut. | June 15, 1993 | Some of the tyvek pouches containing the endo gauge 30 instrument were found to be unsealed, which could compromise the sterility of the device | 43 538 Units Were Distributed. | Nationwide And International.
1993-09-22 | Devices | III | z-782-3 | Ck-mb Flurometric Enzyme Immuno Assay. Recall #z-782-3. | Manufacturer | Baxter Healthcare Corporation Miami Florida. -7- | May 26, 1993 | Phenobarbital conjugate was packed into 10 kits instead of k-mb conjugate | 10 Kits Were Distributed And All Are Accounted For. | Nationwide.